## Slide 1: Ferrite outgassing results G. Cattenoz CWG 01-10-2012 Vincent Baglin, Giulia Lanza, Giuseppe Bregliozzi.

## Slide 2: OUTLINE
Vacuum baseline for LHC
Overview of vacuum validation test
TCTP new features
Thermal outgassing results of ferrites
Estimated outgassing flow of a TCTP collimator
Conclusion

## Slide 3: Vacuum baseline for LHC
“Vacuum requirements for the LHC Collimators”, M. Jimenez, EDMS 428155
UHV cleaned according to CERN specification
Thermal treatment/vacuum firing of material (ex ferrite, CFC, tungsten)
Material to follow bake out at 250°C for 24h00 (48h00 for collimator)
Specific material outgassing flow rate ≈ 10-12 mbar.l/(s.cm2).
Maximum outgassed flow of one collimator 1∙10-7 mbar.l/s

## Slide 4: Overview of vacuum validation test of ferrite
Ferrite type TT2-111R (Supplier: Skyworks) thermally treated under:
Air furnace at 400°C for 24h00 and
Vacuum furnace at 400°C for 24h00.
Vacuum test cycle carried out:
250°C under vacuum for 24h00 + RGA analysis + ultimate pressure record for total N2 equivalent outgassed flow measure by conductance method.
Extended thermal study:
Measure of outgassing flow by accumulation, from R.T. to 250°C.

## Slide 5: Noticeable TCTP features: New RF system
Cross cut of a TCTP collimator:
Total surface of ferrite per collimator = 1600 cm2
BEAM

## Slide 6: Noticeable TCTP features: New tungsten
Tungsten is the constitutive material of the TCTP jaw.
Total surface per TCTP collimator = 2300 cm2
Chosen supplier: MG Sanders Ltd, type DENSAMET® WH9570C .
Outgassing rate results (EDMS 1238270) : 1∙10-12 mbar.l/(s.cm2)
 Match with the LHC vacuum specification

## Slide 7: Ferrite outgassing results
Chart: 1/T [1/K]
| 3.3766999999999998E-3 | 3.3766999999999998E-3 | 3.3766999999999998E-3 | 3.3766999999999998E-3 | 3.3766999999999998E-3 | 3.1730999999999999E-3 | 3.0198E-3 | 3.0016999999999999E-3 | 2.8722999999999999E-3 | 2.8641000000000001E-3 | 2.7766000000000002E-3 | 2.7766000000000002E-3 | 2.6584999999999998E-3 | 2.6584999999999998E-3 | 2.5763000000000001E-3 | 2.5763000000000001E-3 | 2.4865999999999998E-3 | 2.4805000000000001E-3 | 2.3915E-3 | 2.3915E-3 | 2.3086999999999999E-3 | 2.3086999999999999E-3 | 2.2314000000000001E-3 | 2.2314000000000001E-3 | 2.1591000000000002E-3 | 2.1591000000000002E-3 | 2.1591000000000002E-3 | 1.9678999999999999E-3 | 1.9678999999999999E-3
H2 Outgassing from unfired SS - Baked at 300C | 4.29e-12 | 4.08e-12 | 3.74e-12 | 3.87e-12 | 4.29e-12 | 1.45e-11 | 3.68e-11 | 3.88e-11 | 7.71e-11 | 7.88e-11 | 1.33e-10 | 1.39e-10 | 2.5e-10 | 2.61e-10 | 3.64e-10 | 3.68e-10 | 6.27e-10 | 6.53e-10 | 1.08e-09 | 1.26e-09 | 2.54e-09 | 2.18e-09 | 4.14e-09 | 3.76e-09 | 5.28e-09 | 4.18e-09 | 5.41e-09 | 9.95e-09 | 1.05e-08
H2 Outgassing from Fired SS - Baked at 300C | 1.4819e-14 | 1.8935e-14 | 1.5995e-14 | 2.4227e-14 | 1.8464e-14 | 2.8814e-14 | 4.7278e-14 | 7.0682e-14 | 1.2114e-13 | 4.7161e-13 | 4.4456e-12 | 4.0575e-12
Ferrite CMD5005 vac compatible as received | 6.1991e-12 | 6.2947e-11 | 1.5789e-10 | 2.8601e-10 | 4.8179e-10 | 8.0742e-10
Ferrite disk MKI as received | 1.5853e-12 | 2.4912e-12 | 1.2682e-11 | 5.5485e-11 | 2.4346e-10 | 3.4763e-09
Ferrite TT2-111R after TT 400C air and TT 400C vac | 5e-11 | 1.4705e-10 | 1.2471e-10 | 5.137e-10 | 3.1309e-10 | 7.9863e-10 | 8.6243e-10 | 3.5834e-09 | 1.1556e-08 | 1.1412e-08

## Slide 8: Outgassing results – closed up
250°C
200°C
150°C
100°C
50°C
R.T.
Factor 10 gain in outgassing rate for 100°C temperature increase
Chart: 1/T [1/K]
| 3.3766999999999998E-3 | 3.3766999999999998E-3 | 3.3766999999999998E-3 | 3.3766999999999998E-3 | 3.3766999999999998E-3 | 3.1730999999999999E-3 | 3.0198E-3 | 3.0016999999999999E-3 | 2.8722999999999999E-3 | 2.8641000000000001E-3 | 2.7766000000000002E-3 | 2.7766000000000002E-3 | 2.6584999999999998E-3 | 2.6584999999999998E-3 | 2.5763000000000001E-3 | 2.5763000000000001E-3 | 2.4865999999999998E-3 | 2.4805000000000001E-3 | 2.3915E-3 | 2.3915E-3 | 2.3086999999999999E-3 | 2.3086999999999999E-3 | 2.2314000000000001E-3 | 2.2314000000000001E-3 | 2.1591000000000002E-3 | 2.1591000000000002E-3 | 2.1591000000000002E-3 | 1.9678999999999999E-3 | 1.9678999999999999E-3
H2 Outgassing from unfired SS - Baked at 300C | 4.29e-12 | 4.08e-12 | 3.74e-12 | 3.87e-12 | 4.29e-12 | 1.45e-11 | 3.68e-11 | 3.88e-11 | 7.71e-11 | 7.88e-11 | 1.33e-10 | 1.39e-10 | 2.5e-10 | 2.61e-10 | 3.64e-10 | 3.68e-10 | 6.27e-10 | 6.53e-10 | 1.08e-09 | 1.26e-09 | 2.54e-09 | 2.18e-09 | 4.14e-09 | 3.76e-09 | 5.28e-09 | 4.18e-09 | 5.41e-09 | 9.95e-09 | 1.05e-08
H2 Outgassing from Fired SS - Baked at 300C | 1.4819e-14 | 1.8935e-14 | 1.5995e-14 | 2.4227e-14 | 1.8464e-14 | 2.8814e-14 | 4.7278e-14 | 7.0682e-14 | 1.2114e-13 | 4.7161e-13 | 4.4456e-12 | 4.0575e-12
Ferrite CMD5005 vac compatible as received | 6.1991e-12 | 6.2947e-11 | 1.5789e-10 | 2.8601e-10 | 4.8179e-10 | 8.0742e-10
Ferrite disk MKI as received | 1.5853e-12 | 2.4912e-12 | 1.2682e-11 | 5.5485e-11 | 2.4346e-10 | 3.4763e-09
Ferrite TT2-111R after TT 400C air and TT 400C vac | 5e-11 | 1.4705e-10 | 1.2471e-10 | 5.137e-10 | 3.1309e-10 | 7.9863e-10 | 8.6243e-10 | 3.5834e-09 | 1.1556e-08 | 1.1412e-08

## Slide 9: Scaled to TCTP collimator
Estimated outgassing flow for 1 TCTP collimator:
1600 cm2 of ferrite ≈ 8∙10-8 mbar.l/s
2300 cm2 of tungsten ≈ 2∙10-9 mbar.l/s
5000 cm2 of stainless steel ≈ 1∙10-8 mbar.l/s
Total : 9∙10-8 mbar.l/s  Close to LHC vacuum specification limit of 1∙10-7 mbar.l/s

## Slide 10: Outgassing results – Residual Gas Analysis
No contamination visible up to mass 100
RGA scan of ferrite TT2-111R, 48h after bake out, SEM 1400V

## Slide 11: Conclusion
Estimated outgassing flow of one TCTP collimator (9∙10-8 mbar.l/s) is already very close to specified limit (1∙10-7 mbar.l/s)
Should temperature increases from RT to 50°C, outgassed flow is doubled.
Should temperature increases from 50°C to 100°C, outgassed flow increase by factor 10.
RGA analysis shows no contamination
The present solution does not present any safety margin in order to remain within the LHC vacuum specification should the ferrite temperature increase.
VSC proposal: Vacuum test of ferrite with increased thermal treatment temperature.