## Slide 1: CST simulations of the impedance of a new design for the PS wire scanners
F. Caspers, A. Grudiev, E. Métral, B. Salvant
Many thanks to W. Andreazza, R. Calaga and F. Roncarolo (BI)

## Slide 2: Agenda
Context
Main assumptions
Simulation of the wire scanner tank alone
Simulations of the wire scanner with the wire and arm (parking position IN)
Simulations of the wire scanner with the wire and arm (parking position OUT)
Summary

## Slide 3: Context
4 wire scanners are installed in the PS
Beam losses due to small aperture of current wire scanner tank design
 New design for wire scanner in the PS
 See EDMS document 999629
Request for building the new tanks, but impedance had not been checked.
History of broken wires in the SPS due to beam induced losses "Cavity Mode Related Wire Breaking of the SPS Wire Scanners and Loss Measurements of Wire Materials" F. Caspers, B. Dehning, E. Jensen, J. Koopman, J.F. Malo, F. Roncarolo - Proc DIPAC 2003 (pdf) "Wire Measurements on the LHC wire scanner" F. Caspers, T. Kroyer, ABP/RLC meeting (pdf)
2 main risks to this new design  increased impedance (longitudinal and/or transverse) can lead to instabilities
 increased longitudinal impedance can lead to burning the wire

## Slide 4: What can we do to check the impedance?
Theory
Simulations
Measurements
Here we use simulations (CST)
  Time domain  simulates the EM interaction of a line density bunch with a 3D model
     CST Particle Studio Wakefield Solver
  Frequency domain  simulates the modes that can be excited in a 3D model
     CST Microwave Studio Eigenmode Solver

## Slide 5: Main assumptions
Geometrical model
  Hexahedral mesh approximates the structure
  Model can be different from drawing
  Drawing can be different from what is really installed!!!!
Material properties
  In time domain, very coarse fit of the input frequency dependence (see work of Carlo and Lukas Haenichen of TU Darmstadt)
  EM properties of materials at high frequencies are not always measured and can be very different from the model used in CST (very often, they are not part of the specs!!! Also, effect of mechanical and RF damage over time, anisotropies, inhomogeneities, etc.)  See work of Tatiana, Carlo and phase 2 collimation for instance
  Specific issues of modeling a thin wire (see work of Tom Kroyer CERN-AB-Note-2008-018 )
Interaction with the bunch (Particle Studio)
   single pass of a line current density (see work of Carlo and Prof. Vaccaro)
   limited wakelength  limitated to low Q values
   in MWS, perturbation method, well suited to high Q values and no thick dispersive materials

## Slide 6: Agenda
Context
Main assumptions
Simulation of the wire scanner tank alone
Simulations of the wire scanner with the wire and arm (parking position IN)
Simulations of the wire scanner with the wire and arm (parking position OUT)
Summary

## Slide 7: 1. Wire scanner tank alone
Reasons for changing the tank: - improve aperture
- enhance measurement accuracy
CST model after import and processing
STP model exported from CATIA

## Slide 8: Note: import a model from CATIA
Method currently used:
  Ask the model owner to export it in *.STP format
  Open the *.STP file in HFSS (and check the structure…)
  Export in *.SAT format
  Import the *.SAT in CST
  Suppress unnecessary parts and use the CST healing functions
  Fill the simulation space with vacuum and use the separate shape function on the resulting solid to keep the vacuum inside the element and suppress the vacuum around it.
Current issues:
  CERN does not have a direct import license for CATIA or STEP files, hence the need to go through HFSS
  CATIA models are only accessible through SMARTEAM, a shared environment (2-week-training course required)
  Resulting models in CST sometimes present serious issues, and obtaining a consistent model is not always trivial

## Slide 9: Longitudinal wake from CST Particle Studio beta =1, sigma = 2 cm, 2 M mesh cells, wake of ~20 m, indirect testbeams wake integration
0.8 GHz
0.9 GHz
0.95 GHz
1.12 GHz
1.18 GHz
1.45 GHz

## Slide 10: Eigenmodes of the tank
Mode ID | frequency (GHz) | Q | Rs (x,y=0) | Rs/Q | Ploss in W
3 | 0.79 | 5890 | 30 k | 5  | 3 10-7
6 | 0.90 | 4730 | 24 k | 5  | 2 10-10
8 | 0.94 | 6030 | 103 k | 17  | 4 10-11
12 | 1.11 | 4930 | 92 k | 2.8  | 3 10-17
15 | 1.17 | 6730 | 116 k | 17  | 3 10-19
Remarks: Q is obtained with the formula
with W= total stored energy
and P=dissipated power
(linac convention)
(W=1J in eigenmode)
Perturbation method id used to obtain the Q and R for stainless steel.
Very small !!!

## Slide 11: Power losses calculations
If we assume the mode frequency overlaps with one of the beam harmonics (conservative approach)
With the parameters of the LHC nominal beam at ejection of the PS
  nominal bunch charge after splitting q = 18.4 nC (1.15 e11 p/b)
  bunch spacing = 25 ns (worst case scenario)
  smallest nominal RMS bunch length = 30 cm
  Rs is the shunt impedance (linac convention)
  z is the rms bunch length in m

## Slide 12: Why is the dissipated power so small?
Peaks higher than 0.5 GHz are washed out by the large bunch length in the PS
With LHC bunches with small longitudinal emittance, bunch length is around 20 cm,
and dissipated power for mode 1 becomes 0.3 W.
 = 30 cm (PS ejection nominal)
 = 20 cm (PS ejection smaller emittance)
 = 11 cm (LHC injection)
 = 7.5 cm (LHC ejection)
Bunch power spectrum for LHC type bunches of various lengths

## Slide 13: Agenda
Context
Main assumptions
Simulation of the wire scanner tank alone
Simulations of the wire scanner with the wire and arm (parking position IN)
Simulations of the wire scanner with the wire and arm (parking position OUT)
Summary

## Slide 14: 2. Wire scanner tank with wire and mechanical system
CST model after import and processing
STP model exported from CATIA
Ion pump
fork
CST model after import and simplification
Wire at parking position (OUT)
Wire after scan (IN)
Mechanical system

## Slide 15: Tank with wire (IN position) Longitudinal impedance beta =1, sigma = 3 cm, 2 M mesh cells, wake of ~20 m
Additional simulation with Eigenmode solver
(152,000 mesh cells) f0 =292 MHz
Rs= 43 k
Q= 1200
 total power loss in the structure P=800 W for this mode
Fork and casing in stainless steel, wire as PEC wire

## Slide 16: Is this new mode an issue?
 Compare with the losses of the old design
Additional simulation with Eigenmode solver
(156,000 mesh cells) f0 = 292 MHz
Rs= 8.7 k
Q= 730
 total power loss in the structure P=160 W for that mode (new design  800 W)
Fork and casing in stainless steel

## Slide 17: Surface currents (freq domain)
Old design
New design

## Slide 18: Voltage in the wire (1st try…)
Old design
New design
R=1.43 Ohm  P = V2/R = (0.004*18.4)2/1.43 = 4 mW peak in the wire
Charge of the bunch = 1 nC
Rms bunch length: 20 cm
 this method should be crosschecked with the SPS case for which wires actually broke

## Slide 19: Can we damp this mode?

## Slide 20: Trying to damp this 300 MHz mode
Putting some dispersive materials in the structure helps reducing losses due to HOM:
Fritz:
1) Losses are proportional to Rs  “ P = Rs I2 ”
2) Rs/Q is constant whatever the material
3) Both Rs and Q decrease if we create losses in dispersive materials
4) “ P = Rs/Q *Q I2 ”
 create EM losses with dispersive materials
 lower Q for that mode
 fewer losses!
However, of course more losses at low frequencies are expected, and we should be careful with the transverse impedance

## Slide 21: First try to damp that modes with available blocks of ferrite
William says the only possible spots for ferrites are on the top and bottom
They already have 4S60 ferrite plates and they would like to use them if possible.
Possible positions

## Slide 22: Where do we put the ferrite plate
Proposed position for the ferrite

## Slide 23: Longitudinal impedance (real)
Time domain wakefield solver
 Nice damping!
 frequency shifts down
Rms bunch length 20 cm
All materials ss304L except wire
Ferrite 4S60

## Slide 24: Mode properties
 | Simulation type | f0 | Rs | Q | Ploss
No ferrite | eigenmode | 292 MHz | 43 k | 1200 | 800 W
With ferrite | wakefield | 270 MHz | 0.26 k | 21 | 8 W

## Slide 25: Longitudinal impedance (imaginary)
Not too much additional low frequency longitudinal impedance
 maybe not expected
 to be checked
Im(Z/n) = 0.07 to 0.08 

## Slide 26: Horizontal “dipolar” impedance (LF)
No symmetry…
We displace the beam by 10 mm towards the engine and look for the slope of the horizontal impedance (first approximation)
Imaginary LF horizontal “dipolar” impedance =(20-8)/0.01= 1.2 k/m

## Slide 27: Vertical dipolar impedance (LF)
There is a symmetry in the vertical plane
Imaginary LF vertical dipolar impedance =13/0.01= 1.3 k/m

## Slide 28: Agenda
Context
Main assumptions
Simulation of the wire scanner tank alone
Simulations of the wire scanner with the wire and arm (parking position IN)
Simulations of the wire scanner with the wire and arm (parking position OUT)
Summary

## Slide 29: Tank with wire (OUT position)
Wire position |  | Simulation type | f0 | Rs | Q | Ploss
OUT | No ferrite | eigenmode | 264 MHz | 2.6 k | 853 | 89 W
OUT | With ferrite | wakefield | 256 MHz | 0.06 k | 30 | 2 W
IN | No ferrite | eigenmode | 292 MHz | 43 k | 1200 | 800 W
IN | With ferrite | wakefield | 270 MHz | 0.26 k | 21 | 8 W
No other harmful mode detected
Mode due to the wire and fork
Longitudinal: Z/n=0.04 Ohm
horizontal: 0.9 kOhm/m
Vertical: 0.5 kOhm/m

## Slide 30: Summary
Need to check the new wire scanner design for the PS
From these simulations, new design seems to create more longitudinal impedance and more losses.
From the figures obtained, this increase does not appear huge
What should we do?
 recommendation for short term:
build like this and be ready to install ferrites if needed.
 recommendation for longer term:
Compare these figures with simulation of the SPS wire that broke. Perform wire measurements when the tank is built.
optimise the position, shape and material of the ferrites

## Slide 31: Thank you for your attention

## Slide 32: Mode 3
E field
H field

## Slide 33: Mode 6

## Slide 34: Longitudinal impedance

## Slide 35: Horizontal dipolar impedance at low frequency
LF simulation parameters:
40,000 mesh cells
Rms bunch length=20cm

## Slide 36: Ez Field patterns of this mode

## Slide 37: |H| field pattern of this mode