## Slide 1: BSRT UPDATE LMC 02-OCT-2012 F.Roncarolo on behalf of the BSRT team W.Andreazza, E.Bravin, A.Boccardi, J-J.Gras, A.Goldblatt, M.Hamani, T.Lefevre, R.Jones, A.Nosych, G.Trad, R.Veness, M.Wendt A.Fisher EN-MME – S.Sgobba and team, Main Workshop, Design Office E.Metral, B.Salvant, F.Caspers, A.Grudiev, C.Vollinger G.Lanza and Team

## Slide 2: Recap - TS#3 Actions
BSRT mirrors before TS#3: silicon bulk + dielectric coating
After B2 mirror removal end of August (coating blistered) and suspect of same damage on B1 mirror, three options for TS#3:
  Spare mirror: Silicon bulk + dielectric coating
    Option dropped, good choice after B1 mirror removal (coating blistered)
  Spare mirror after coating removal and polishing  pure silicon mirror
    Installed on B2
  New mirror: glass bulk + metallic (Al) coating
    Installed on B1
On both B1 and B2:
  Failsafe mirror supports
  Inconel mirror clamps
  Ferrites with higher Curie temperature – increase ferrite volume
See LMC 14 Sep 2012

## Slide 3: B1 Mirror removal
Evidence of blistering
Mirror out of place
Decided not to install an identical mirror

## Slide 4: B1 refurbishment during TS#3
Mirror: glass bulk+ metallic (Al) coating
Imaging optics changed from focusing mirrors to focusing lenses
  Change already scheduled in advance, independently from August’s systems failure
  Alignment with calibration laser as done with old optics
OLD: 2 focusing + 8 folding mirrors
NEW: 2 focusing lenses + 2 folding mirrors
A.Goldblatt, A.Fisher

## Slide 5: B1 performance @ re-start
Steering to find light spot very similar to what expected after laser alignment
Amount of light ~ as expected, ~ = dielectric coating (maybe even better, to be characterized)
Preliminary calibration w.r.t. WS (next slides)

## Slide 6: BSRT vs WS B1 HOR 4 TeV
BSRT Correction in quadrature: 0.21 mm (was 0.38 mm)
G.Trad

## Slide 7: BSRT vs WS B1 VER 4TeV
BSRT Correction in quadrature: 0.23 mm (was 0.35 mm)
G.Trad

## Slide 8: B2 refurbishment during TS#3
Mirror: polished Silicon bulk, no coating
Imaging optics (focusing mirrors) unchanged
  Re-alignment with calibration laser
  Imaging of calibration target
Optical table top view
Illuminated target
Target calibration line
Extraction mirror (vacuum) + steering mirror (air, first mirror on the table)
BSRT Camera
Imaging calibration target as before TS#3
 no evidence of image distortion on all mirrors apart from Extraction + Steering mirrors

## Slide 9: B2 observations @ re-start
Steering to find light spot very similar to what expected after laser alignment
Amount of light ~ as expected after removing coating (~ factor 2 less w.r.t. coated mirror)
Bad surprise looking at image quality
450 GeV
4 TeV

## Slide 10: B2 Scan mirror IN  OUT
1
0
1
2
3
4
Conclusion: Moving mirror out
 loose image ‘core’, ‘fringes’ remain
Proton beam
Sync. Light
BSRT mirror
MIRROR IN
4 TeV
N.B.: tried same exercise on B1  no evidence of distorted images…
Proton beam
Sync. Light
BSRT mirror
MIRROR OUT

## Slide 11: B2 Selecting different wavelengths
All wavelengths
Blue
Red
Green
Red
Larger radiation wavelength  larger ‘fringes’ distance  compatible with diffraction
4 TeV

## Slide 12: B2 Horizontal Closed Orbit Bumps
NO bump
+2mm Bump
+4mm Bump
-4mm Bump
0
Proton beam
Sync. Light
BSRT mirror
Bump -
Bump +
70 mm
In all these cases: camera gain to have maximum ‘fringes’ visibility
Bump +  weight of ‘fringes’ w.r.t. ‘core’ decreases
4 TeV

## Slide 13: B2 Horizontal Closed Orbit Bumps
+4mm bump, low gain
NO bump, low gain
N.B. :
absolute position on screen depends on steering (not the same for these two acq.)
Verified that steering ( different path inside telescope) doesn’t change ‘fringes’ shape and weight
Anyway: emittances from Gaussian fits give unphysical large emittance
4 TeV
Setting camera gain as in operational mode (no saturation)

## Slide 14: Temperatures Probes
Flange 1
Flange 2
View Port 1
View Port 2
Bellow
See Measurements next slide

## Slide 15: Temperatures
B1 heating ~= B2 heating
Bellow heats more and faster than flange (== radiation more than conduction?)
Mirror OUT
Mirror IN
Mirror Falling
Before TS#3 (‘bellow’ probe B1 not at same location as B2)
After TS#3
‘bellow’ probe B1 displaced to be = B2
Added 2 probes on flange
BEAM 1
BEAM 2
1380 b
1380 b
Bellow
Bellow
Flange
Flange
View Port
View Port
1380 b
1380 b
1380 b
1380 b
LAMP ON

## Slide 16: Vacuum
5.L4  BSRT B2 , baked out during TS#3 ….

## Slide 17: Summary of Observations after TS#3
B1 ok, need to verify if survives to thermal cycles
  New optics base on lenses seems better for overall resolution and accuracy, to be confirmed
B2 image ‘distortion’ to be understood
  edge effect?
  Bad surface polishing?
  Silicon not suitable for imaging?
  Optics after extraction mirror misalignment?
No clear evidence of less heating w.r.t before TS#3 for both B1 and B2
No clear evidence of B1 (metallic coating) heating more than B2
Both B1 and B2 AGM operational despite bad B2 imaging. Calibration checks on-going, need to verify calibration stability over time

## Slide 18: Plans - Proposals
  ASAP:
    access 3 hours  re-check alignment B2 with calibration laser  exclude (or not) any source of uncertainty other than the extraction mirror itself
    Test imaging system based on pure silicon mirror in the lab (on-going)
  Before MD3 (if possible)
    B2 Bumps (450 GeV and/or 4 TeV) + move mirror out with bump >= 4mm IN  assess it’s only an edge effect or not
  MD3
    B1: Precise Calibration w.r.t WS
    B2: points above if not possible before
    If enough time: B1&B2, characterize heating with high intensity beams at 450 GeV
      Wait between injections to cope with thermal inertia and establish whether there is a threshold effect or not
  TS#4 (xstmas)
    depending on possible new observations/understanding, e.g. will B1 mirror survive thermal cycles?
      Open B2 and replace mirror with copy of B1 (glass + metallic coating)
    If new B1 optics full validated
      Replace also B2 optics with focusing lenses based system
  LS1: depending on end of the run experience + improved RF and thermo-mechanical simulations
    Redesign mechanics to minimize RF coupling
    Active cooling
    …

## Slide 19: SPARE